| Written/updated by |  | Date | Effective date/reference |
| --- | --- | --- | --- |
| Petra Kristian, Brittanie Fullerton |  |  |  |
| Reviewed by | Signature | Date |  |
| Approved by | Signature | Date |  |

# Purpose

The purpose of the formative evaluation test protocol is to ensure that the formative evaluation test is carried out as intended and that the test can be repeated.

# Scope

This document includes all information about the formative evaluation test for PhenoMap in full detail.

# Contents

1	Purpose	2

2	Scope	2

3	Contents	2

4	Terms and definitions	3

5	Test purpose	3

6	Method	4

7	Test object	4

8	Test setup and material	4

9	Test participants	5

10	Consent procedure	6

11	Participant training	6

12	Test sequence and duration	6

13	Selection of tasks	7

14	Tasks overview	7

15	Evaluation activities	8

15.1	Use scenarios	8

15.2	Interview	8

16	Methods for data collection	8

17	Data analyses and reporting	8

17.1	Data collection	8

17.2	Data analysis	9

17.3	Deviations	9

18	Change history	9

# Terms and definitions

| Term | Definition |
| --- | --- |
| User interface | Means by which the user and the medical device interact. |
| User interface evaluation | Process by which the manufacturer explorers or assesses the user interaction with the user interface |
| Usability engineering / Human factors | Application of knowledge about human behaviour, abilities, limitations, and other characteristics to the design of medical devices (including software), systems, and tasks to achieve adequate Usability |
| Formative evaluation / Test | User Interface evaluation conducted during the user interface development. |
| Hazard-related use scenarios | Use Scenario that could lead to a hazardous situation or harm |
| Critical task | Tasks that, if performed incorrectly or not performed at all, would or could cause serious harm to the patient or user. |
| Test object | The device or the prototype with which the test will be conducted |
| Test participant | Person that participates in the test and corresponds to the intended user group as much as possible. |
| IFU | Abbreviation of “Instructions for use” (user’s manual) |

# Test purpose

The goal of this evaluation is to:

A. explore user interface strengths and weaknesses,

B. identify potential use errors that would or could result in harm to the patient or user,

C. find potential areas for improvement of the design, in regard with the use of PhenoMap.

# Method

The evaluation will be conducted as a usability test (simulated use test + interview) within simulated-use approach. This method allows the observation of potential misuse and clustering misinterpretability cases. The protocol is as follows: (1) The participant will be briefed on the device purpose; (2) The participant will be presented with various possible new phenotypes generated by the clustering algorithm and patient data following each cluster; (3) The participant will be encouraged to verbalize their thought process and justify their interpretation of classifying and identifying patient clusters and outliers. The evaluation will be conducted with a representative clinician and has a duration of 40 minutes.

# Test object

A functional software prototype featuring:

- Backend: PAM clustering and Random Forest outcome models.

- GUI: PCA and UMAP/t-SNE/HDBScan scatter plots, biomarker (feature importance) heatmaps, and GOSE prediction displays.

- IFU: Preliminary guide for interpreting phenotypes and outlier classifications.

# Test setup and material

The following table describes the parts of the user interface - with their associated version number- that will be tested.

| # | Component(s) or device tested | Version |
| --- | --- | --- |
| 1 | Preprocessing and PCA filtering | {1.0} |
| 2 | Discovery interface (UMAP/t-SNE/HDBScan & PAM clustering) | {1.0} |
| 3 | Outcome validation (GOS-E and Random Forest) | {1.0} |
| 4 | Interpretability (SHAP feature importance) | {1.0} |

The following table describes the materials and supplies that will be used to conduct the testing.

| # | Material(s), supplies used to conduct the test | Quantity |
| --- | --- | --- |
| 1 | p-TBI dataset | {51 patients} |
| 2 | Instruction for use (IFU) | {1} |

The following table describes the equipment that will be used to conduct the testing.

| Equipment used to conduct test |  |
| --- | --- |
| 1 | Laptop and monitor |
| 2 | Moderation guide and prompts |

# Test environment

The evaluation was conducted via Zoom and the participant accessed the software through screen sharing.

# Test participants

An expert clinician, Dr. Vaccari, participated as a representative of the user group. This selection facilitated a technical validation of the sub-phenotype logic and ensured the diagnostic labels accurately reflected complex pediatric TBI trajectories.

# Participant training

No training is given in advance of the use of PhenoMap.

# Test sequence and duration

Test sequence with estimated duration:

| Test sequence | Estimated duration [min] |
| --- | --- |
| Introduction | 5 min |
| Confirm testing start | 2 min |
| US 1 – Data Exploration (Reviewing PCA, clusters, and features) | 15 min |
| US 2 – Clinical Validation (Checking SHAP importance and GOSE outcomes | 10 min |
| Post-hoc Interview (Discussion on sub-phenotype utility) | 5 min |
| Completion | 3 min |

Total estimated time for each test is 40 minutes.

# Selection of tasks

The formative evaluation tests critical tasks related to the digital workflow of identifying p-TBI phenotypes. Participants perform tasks focused on data interpretation, such as grouping patients by biomarker levels and validating through the algorithm’s recovery and supervised predictions. The test also covers feature selection to see how adding or removing certain biomarkers changes the results, helping find design improvements for clinical use.

# Tasks overview

| Test Section | Task | Critical Task ID / HRUS ID |
| --- | --- | --- |
| US 1 – Data Exploration (Reviewing PCA, clusters, and features) | Select features | CT1.1 |
|  | Adjust clustering | CT1.2 |
|  | Visualize clusters | T1.3 |
| US 2 – Performance of {blood glucose measurement} | Interpret SHAP importance | CT2.1 |
|  | Validate outcomes | CT2.2 |
|  | Identify outliers | T2.3 |

# Evaluation activities

## Use scenarios

The moderator will screen share the clustering, biomarker data, and other visuals for the clinician to analyze in real-time. There will be a discussion and advice session from the clinician.

## Interview

Following each scenario, the moderator will conduct a brief interview to capture subjective feedback on the prototype's interpretability. The strategy includes the following questions:

- How easy was it to distinguish between the sub-phenotypes?

- How could the visualization of biomarker importance (SHAP) be made clearer?

- Was there any data on the dashboard (e.g., specific PCA dimensions) that you did not need?

- Were you missing any specific clinical metrics, such as a direct comparison to standard GCS scores?

# Methods for data collection

During testing, the following objective performance data and subjective feedback will be collected:

- Findings (i.e., visualization of clustering algorithms, difficulties, etc.)

- Participants’ subjective assessments regarding the system

- Moderator observations

The moderator will make note of key data and observations.

# Data analyses and reporting

## Data collection

The collected data from notes taken during the study will be analysed.

## Data analysis

The data collection will be analysed to identify trends and patterns regarding use-related risks and usability.

The data analysis contains the following step:

Identify the conditions and causes that led to observed use errors, i.e., a detailed root cause analysis for use errors by means of follow-up interviews, participant comments and judgement, and observer notes.

# Change history

| Revision | Ref | Description |
| --- | --- | --- |
| 1.0 | Whole document | First release |