Short SaMD Design Project Report:

- Final SaMD technology title:

  - PhenoMap

- Final intended use statements:

  - PhenoMap is a Software as a Medical Device (SaMD) intended for use by pediatric intensive care clinicians and neurosurgeons to identify and visualize objective biological sub-phenotypes in pediatric patients. This technology specifically analyzes temporal dynamics of neuro-specific inflammasomes alongside CNS biomarkers (at 0h and 24h) to group patients into data-driven recovery trajectories that may be invisible to traditional clinical assessments like the Glasgow Coma Scale (GCS). It is designed to be utilized within the first 24 hours following a traumatic brain injury (TBI), exclusively within tertiary trauma centers, emergency rooms, and pediatric intensive care units (PICUs). Clinicians will access the technology via a MATLAB-based graphical user interface (GUI) to review cluster assignments and feature importance to inform personalized neuro-protective strategies.

- Proposed risk-based device classification:

  - Class II.

- Proposed regulatory pathway:

  - De Novo classification request.

- Overview of final ML approach:

  - This SaMD uses an unsupervised to supervised pipeline.

  - Dimensionality Reduction: Principal Component Analysis (PCA) is used to reduce high-dimensional biomarker data (inflammasomes and clinical information), retaining components explaining >80% of total variance.

  - Clustering: Partitioning Around Medoids (PAM) with k=3 (compared to k=6 (optimal)) was implemented on PCA scores to ensure robust, deterministic phenotype assignment.

  - Interpretability: A Random Forest “One-vs-Rest” analysis identifies top feature drivers (e.g., NFL, Weight, Age) to provide clinical transparency for each cluster assignment.

- Overview of final ML performance metrics:

  - Supervised Validation: A Random Forest classifier was trained on the cluster assignments to assess phenotype separability. The model achieved a low Out-of-Bag (OOB) error and high accuracy, confirming that the identified clusters are defined by distinct, non-overlapping biological signatures.

  - Prognostic Significance: The primary performance metric was the differentiation of 12-month recovery outcomes. The k=3 model achieved a significant p = 0.0286 (Kruskal-Wallis) for GOS-E Peds scores at 6 – 12 months. This is then compared to the k=6 (optimal) model, which achieved a p = 0.0283 for GOS-E scores at 2 – 6 weeks.

  - Clustering Quality: Achieved a positive average Silhouette Coefficient of 0.38 for both clustering assignments (k=3 vs. k=6 (optimal)).

  - Reliability: The model demonstrated 100% reproducibility in phenotype assignment using a deterministic PCA-based coordinate system.

- Overview of additional related information (as needed):

  - Human-in-the-Loop (HITL): The SaMD includes a “Phenotype Fingerprint” heatmap and “Explain Plots” (Pareto) to allow clinicians to manually verify the AI's logic against raw patient lab values.